## Extended
Name of Test (linked to IFU) | Manufacturer (linked to page) | Type of test | Sensitivity | Specificity | Description | Phase of development
SGTi-flex COVID-19 IgG | Sugentech Inc. | RST | 41.2% (0-7 days), 91.7% (7-14 days), 98.6% (15+ days) | 99.15% | The SGTi-flex COVID-19 IgG is a colloidal gold based lateral flow assay, specific for IgG to SARS-CoV-2. Samples should only be tested from individuals 15 days or more post-symptom onset. The test strip is within a cassette, and the results can be read by eye within about 15 minutes. The assay detects IgG specific to the SARS-CoV-2 nucleocapsid and RBD of the spike protein. This test can use human serum, plasma, or whole blood samples--including fingerstick whole blood samples. Sensitivity was determined from 185 positive clinical samples and stratified by time post symptom onset, and specificity from 234 negative clinical samples. The test was also validated by the NCI, with a sensitivity of 96.7%, and a specificity of 100%. The April 2021 updates includes language to clarify that the test should not be used to measure vaccine responses, and that performance of the test has not been established in individuals that have received a COVID-19 vaccine. They also state that the serology test has not been validated for all variants circulating. | Received FDA EUA, CE mark, Korea MFDS Product license, and Brazil ANVISA certified
Rapid COVID-19 IgM/IgG Combo Test Kit | Megna Health, Inc. | RST | IgM: 73.81%, IgG, 73.81, Combined: 90.5% | IgM: 99.7%, IgG:99.3%, Combined: 98.9% | This is a lateral flow assay rapid serology test that tests for IgM and IgG to SARS-CoV-2 nucleocapsid protein. Sensitivity was determined from 126 positive clinical samples, and specificity was determined from 285 negative samples. They present a combined (IgM and IgG) sensitivity, but also stratify by isotype and time since symptom onset. They also present independent NCI validation results, with combined sensitivity of 100% and specificity of 95%. The April 2021 update includes the addition of using fingerstick whole blood samples. In addition, they added language to clarify that the test should not be used to measure vaccine responses, and that performance of the test has not been established in individuals that have received a COVID-19 vaccine. They also state that the serology test has not been validated for all variants circulating. | Received FDA EUA
QUANTA Flash SARS-CoV-2 IgG | Inova Diagnostics, Inc. | ChLIA | 66.7% (0-7 days), 61.5% (7-14 days), and 100% (15+ days) | 99.9% | The QUANTA Flash SARS-CoV-2 IgG test is a chemiluminescent immunoassay for the qualitative and semi-quantitative detection of IgG in patient serum in patients 15 days or more post-symptom onset. The test detects IgG specific to the nucleocapsid and spike antigens of SARS-CoV-2. The test should be run using the BioFlash instrument platform, which allows for antigen-coated paramagnetic beads to mix with patient serum, and antibody levels are then correlated to relative light units. Sensitivity was determined using 51 positive clinical samples, stratified by time post rRT-PCR positive. Specificity was determined using 1028 negative clinical samples (528 of which had underlying conditions). The test should not be used to measure vaccine responses, and that performance of the test has not been established in individuals that have received a COVID-19 vaccine. They also state that the serology test has not been validated for all variants circulating. | Received EUA by the FDA
Simoa Semi-Quantitative SARS-CoV-2 IgG Antibody Test | Quanterix Corporation | paramagnetic microbead-based sandwich ELISA | 45.2% (0-7 days), 87.5% (7-14 days), and 100% (15+ days) | 99.2%% | The Simoa Semi-Quantitative SARS-CoV-2 IgG Antibody Test uses paramagnetic microbead-based sandwich ELISA coupled with single molecule array (Simoa) technology to detect antibodies with high sensitivity. Briefly, the assay depends on a magentic bead binding an antibody of interest, a complex which is then bound by a detection antibody. This is flowed over microwells that each hold one bead, and then the signal may be detected to determine if antibodies of interest are present. The assay is semi-quantitative, detecting IgG antibodies against the spike protein. Time to perform 96 tests is 2 hours and 45 minutes, and the test requires the Simoa HD-X Analyzer platform. Sensitivity was determined using 75 positive clinical samples stratified by time post-PCR positivity, and specificity using 496 negative samples. The April 2021 update includes data on electromagnetic compatibility (EMC) testing for the Quanterix HD-X platform, and importantly, states that this test should not be used in patients taking biotin supplements. Biotin may interfere with the test and produce false results. Proteins, including bilirubin and rheumatoid factor, may also decrease positive signal. It also includes language stating that the test should not be used to measure vaccine responses, and that performance of the test has not been established in individuals that have received a COVID-19 vaccine. They also state that the serology test has not been validated for all variants circulating. | Received EUA by the FDA
MidaSpot COVID-19 Antibody Combo Detection Kit | Nirmidas Biotech, Inc. | RST | 76.9% (7-14 days), 100% (15+ days) | 100% | The MidaSpot COVID-19 Antibody Combo Detection Kit is an RST that can use serum, plasma, and fingerstick whole blood samples to detect IgM and IgG antibodies. Serum and plasma samples are only authorized for use in CLIA labs, while fingerstick samples are authorized in POC settings with a CLIA Certificate of Waiver, Certificate of Compliance, or Certificate of Accreditation. Combined IgG/IgM sensitivity was determined using 48 positive clinical samples stratified by time post symptom onset, and specificity using 100 negative samples. The April 2021 update includes shelf life stability (6 months when stored 4C-25C). It also includes language stating that the test should not be used to measure vaccine responses, and that performance of the test has not been established in individuals that have received a COVID-19 vaccine. They also state that the serology test has not been validated for all variants circulating. | Received EUA by the FDA
COVID-19 Self-Collected Antibody Test System | Symbiotica, Inc. | Home collection, ELISA | 100% (8-14 days), 100% (15+ days) | 98% | The COVID-19 Self-Collected Antibody Test System is a fingerstick self-collection kit. The blood sample (a dried blood spot-DBS) is sent via USPS to Symbiotica, Inc labs, where it is tested in an ELISA based assay. The fingerstick can be self-collected in individuals 18 and older, and in individuals 5 and older with the help of an adult. The ELISA itself is based off of the EUROIMMUN Anti-SARS-CoV-2 Assay, which tests for IgG specific to SARS-CoV-2. Self collection is monitored through the InTeleLabs telehealth system, with results reported via the elicity platform. Sensitivity was determined from 32 samples (from symptomatic patients) tested positive by PCR at least 8 days but no more than 45 days prior to sampling. Only 4 samples were tested from days 8-14. Specificity was determined using 51 negative samples. They also performed sample stability studies in high and low temperatures, and usability studies in adults and children. Sample testing is limited to Symbiotica, Inc., located at 1350 Burton Drive, Ste 210 Vacaville, CA 95687. | Received FDA EUA
Dimension EXL SARS‑CoV‑2 IgG (CV2G) | Siemens Healthcare Diagnostics Inc. | ChLIA | 61.1% (0-7 days), 92.9% (7-14 days), 100% (15+ days) | 100% | The Dimension EXL SARS‑CoV‑2 IgG (CV2G) assay is a chemiluminescent immunoassay that can provide qualitative and semiquantitative data on IgG levels specific to SARS-CoV-2. This test should be run on the Dimension EXL platform. The currently published IFU is shared with this assay and the Dimension VIsta SARS-CoV-2 IgG assay. Sensitivity was determined from 245 positive clinical samples (stratified by time post-positive PCR), and specificity from 1509 negative clinical samples. They also performed precision studies and seroconversion studies, published in the IFU. The March 2021 updates include information regarding vaccination and variants. Specifically, this test has not be validated in individuals who have been vaccinated. The significance of a positive or negative test results should not be used to interpret the efficacy of protection provided by a vaccine. Performance was not established with all currently circulating variants, and performance may change depending on variants present. | Received FDA EUA.
Dimension Vista SARS‑CoV‑2 IgG (COV2G) | Siemens Healthcare Diagnostics Inc. | ChLIA | 61.1% (0-7 days), 92.9% (7-14 days), 100% (15+ days) | 100% | The Dimension Vista SARS‑CoV‑2 IgG (COV2G) is a chemiluminescent immunoassay that can provide qualitative or semi-quantitative data on patient IgG levels specific to SARS-CoV-2. This CLIA uses luminescent oxygen channeling immunoassay technology to capture and detect antibodies bound to beads. The target antigen is spike protein RBD. The test should be run on the Dimension Vista platform. Sensitivity was determined from 245 positive clinical samples (stratified by time post-positive PCR), and specificity from 1509 negative clinical samples. They also performed precision studies and seroconversion studies, published in the IFU. The March 2021 updates include information regarding vaccination and variants. Specifically, this test has not be validated in individuals who have been vaccinated. The significance of a positive or negative test results should not be used to interpret the efficacy of protection provided by a vaccine. Performance was not established with all currently circulating variants, and performance may change depending on variants present. | Received FDA EUA.
VITROS Immunodiagnostic Products Anti-SARS-CoV-2 Total Reagent Pack | Ortho Clinical Diagnostics | ChLIA | 80% (0-8 days), 100% (8+ days) | 100% | The March 2021 update includes information relevant to individuals who have been vaccinated. Specifically, this test has not be validated in individuals who have been vaccinated. The significance of a positive or negative test results should not be used to interpret the efficacy of protection provided by a vaccine. The February 2021 update includes language specific to results. "Positive" results should now be considered "reactive," and negative results, "non-reactive." In addition, the calibrator and controls must be shipped and stored at a temperature of -20C or lower (frozen).The October update includes language on false positives, clearly specifying that a positive result does not necessarily indicate prior infection, and that a second test may be used to confirm. Specifically, they state "Due to the risk of false positive results, confirmation of positive results should be considered – using a second, different antibody assay that detects the same type of antibodies." This test is a proprietary ELISA, and detects total IgM and IgG, but does not discern between the two. The target antigen is SARS-CoV-2 spike protein. Sensitivity was 83% when tested in 36 samples known positive, and sensitivity was 100% out of 400 known SARS-CoV-2 negative samples. Sensitivity increases as day from symptom onset increases. This must be used on the platform VITROS® XT 7600 Integrated System, the VITROS® 3600 Immunodiagnostic System, the VITROS® 5600 Integrated System and VITROS® ECi/ECiQ Immunodiagnostic systems. | Received EUA by the FDA
SCoV-2 Detect IgM ELISA | InBios International, Inc. | ELISA | 92.5% | 98.5% | This kit specifically tests for IgM present in patient serum. Samples should only be tested from individuals that are 7-64 days post symptom onset. SARS-CoV-2 antibody negative samples collected 7 days or more post symptom onset should be reflexed to a test that detects and reports SARS-CoV-2 IgG.The target antigen is not stated. Samples for testing should only be taken between 7 days to 64 days post symptom onset. Samples taken before 12 days post-symptom onset that test negative should be re-tested using a second, molecular method. Samples, taken after 12 days post symptom onset, that are negative should also be tested for IgG. Sensitivity was determined from 120 positive clinical samples, and specificity from 95 negative samples. Sensitivity presented is overall sensitivity, but the company does list varied sensitivity stratified by time post symptom onset. In general, the greater number of days post symptom onset, the more sensitive the assay. | Received EUA
Access SARS-CoV-2 IgG II | Beckman Coulter, Inc. | ELISA | 81.8% (0-7 days), 95.8% (8-14 days), 98.9% (15+ days) | 100% | The Access SARS-CoV-2 II IgG assay is for both qualitative and semi-quantitative detection of IgG antibodies to the receptor binding domain of the spike protein. The assay is a two-step ELISA, and should be run on an Access Family platform. For quantitative detection, the numerical result in arbitrary units (AU) can be from 2.00-450 AU/mL1, and a qualitative result reported as non-reactive (<10 AU/mL) or reactive (greater than or equal to 10 AU/mL) to SARS-CoV-2. Sensitivity was determined using 126 positive clinical samples, including 91 from 15 days or more post symptom onset. Specificity was determined using 93 negative samples. | Received EUA by the FDA. Testing is limited to laboratories designated by Adaptive Biotechnologies Corporation that includes the Adaptive Biotechnologies Lab located at 1551 Eastlake Ave E Ste 200, Seattle, Washington, which is also certified under the Clinical Laboratory Improvements Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meets requirements to perform highcomplexity tests.
T-Detect COVID Test | Adaptive Biotechnologies Corporation | multiplex PCR, NGS | 97.1% | 98.9% | The T-Detect COVID Test is a combination of multiplex PCR and next generation sequencing to identify the expression of the T-cell receptor β (TCRβ) genes. In other words, it assays the activity and expression levels of T-cells (part of the cellular immune response) that are specific to SARS-CoV-2. Specimens should only be tested from individuals that are 15 days or more post-symptom onset, and should be venous whole blood samples. The test depends on an algorithm that identifies rearranged TCRβ sequences from human genomic DNA that indicate an adaptive response to SARS-CoV-2 infection. Sensitivity and specificity were determined using samples that had been tested by PCR (not a T-cell specific assay). Sensitivity was determined using 38 samples 15 days or more post symptom onset, and specificity using 99 negative clinical samples. | Received EUA by the FDA. Testing is limited to laboratories designated by Adaptive Biotechnologies Corporation that includes the Adaptive Biotechnologies Lab located at 1551 Eastlake Ave E Ste 200, Seattle, Washington, which is also certified under the Clinical Laboratory Improvements Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meets requirements to perform highcomplexity tests.
AdviseDx SARS-CoV-2 IgG II | Abbott Laboratories Inc. | CMIA | ARCHITECT: 49.33% (0-7 days), 82.6% (7-14 days), 98.1% (15+ days); Alinity I: 49.33% (0-7 days), 80.43% (7-14 days), 98.1% (15+ days) | 99.55% (both ARCHITECT and Alinity I) | The AdviseDX SARS-COV-2 IgG II assay is a qualitative and semi-quantitative assay to detect IgG antibodies to SARS-CoV-2 in human serum. This assay can be run on the ARCHITECT or AlinityI systems. The target antigen is the RBD region of the spike protein. Magnetic, RBD-coated particles are incubated with patient serum, and any RBD-specific IgG antibodies present will bind the magnetic particles. The particles are then washed, and a chemiluminescent label is incubated, producing a chemiluminescent readout of IgG levels. The limit of quantitation is 22 AU/mL. Sensitivity was calculated using both days post symptom onset, and days post-positive PCR. The data presented here are post symptom onset. Sensitivity was determined using 220 positive clinical samples (stratified by time), and specificity using 2,008 negative clinical samples. | Received EUA by the FDA
IDS SARS-CoV-2 IgG | Immunodiagnostic Systems Ltd | ChLIA | 60.9% (0-7 days), 80.% (7-14 days), 97.6% (15+ days) | 99.6% | The IDS SARS-CoV-2 IgG assay is a qualitative assay for detection of IgG antibodies specific to the nucleocapsid and spike proteins of SARS-CoV-2. The chemiluminescent immunoassay provides a light-based readout of quantity of antibodies, and should be run on the IDS-iSYS MultiDiscipline Automated System. Sensitivity was determined using 141 positive clinical samples, stratified by time post PCR positivity. Specificity was determined using 554 negative clinical samples. | Received EUA by the FDA
UBI SARS-CoV-2 ELISA | United Biomedical, Inc. | ELISA | 38.5% (0-7 days), 79.5% (7-14 days), 100% (15+ days) | 100% | The UBI® SARS-CoV-2 ELISA is an immunoassay that employs synthetic peptides derived from the Matrix(M), Spike(S) and Nucleocapsid(N) proteins of SARS-CoV-2 for the detection of antibodies to SARSCoV-2 in human sera or plasma. A blue color develops in proportion to the amount of SARS-CoV-2-specific IgG antibodies present, if any, in most settings. Absorbance of each well is measured within 15 minutes at 450 nm by using a microplate reader such as a VersaMax™ by Molecular Devices® or equivalent. Sensitivity and specificity were determined from 53 positive and 154 negative clinical samples. | Received FDA EUA.
EliA SARS-CoV-2-Sp1 IgG Test | Phadia AB | Fluoroenzyme Immunoassay | 15.4% (0-7 days), 58.2% (7-14 days), 97.6% (15+ days) | 99.4% | TheEliASARS-CoV-2-Sp1IgGtestisafluoroenzymeimmunoassay intendedforqualitative and semi-quantitative detection of IgG antibodies to SARS-CoV-2 spike 1 in human serum and plasma using the Phadia 250 instrument. The EliA SARS-CoV-2-Sp1 IgG Wells are coated with recombinant SARS-CoV-2 spike 1 protein. If present in the patient's specimen, antibodies to SARS-CoV-2 spike 1 protein bind to their specific antigen. Sensitivity and specificity were determined from 163 positive and 330 negative clinical samples. | Received FDA EUA and the CE mark
RapCov Rapid COVID-19 Test | ADVAITE, Inc. | RST | 90% | 95.2% | The RapCov™ Rapid COVID-19 Test is a lateral flow immunoassay intended for the qualitative detection of IgG antibodies to the SARS-CoV-2 virus in human fingerstick whole blood samples. The antigen (Recombinant SARS-CoV-2 Nucleocapsid Protein) is conjugated to colloidal gold nanoparticles and the antigen-gold conjugate moves to the test line area and attaches to the specific IgG antibodies to SARS-CoV-2. Sensitivity and specificity were determined from 30 positive and 104 negative clinical samples. | Received FDA EUA.
RightSign COVID-19 IgG/IgM Rapid Test Cassette | Hangzhou Laihe Biotech Co., Ltd | RST | 100% (IgM, 0-6 days), 0% (IgG, 0-6 days); 85.7% (IgM, 7-14 days), 76% (IgG, 7-14 days); 99.25% (IgM, 14+ days), 98.5% (IgG, 14+ days) | 99.43% | The December 2 update includes the distributor Unisources Group LLC. This test is a lateral flow assay/colloidal gold method to detect IgG and IgM antibodies. The target antigen is the S1 region of the spike protein. They used 178 positive samples to determine sensitivity, stratified by day since symptom onset. They determined specificity from 349 negative samples, 239 of which were from patients with non-SARS-CoV-2 coronavirus infections. The test was also independently validated by the NCI, and found 96.7% IgM sensitivity, 100% IgG sensitivity, and 98.8% specificity. | Received EUA by the FDA
ACON SARS-CoV-2 IgG/IgM Rapid Test | ACON Laboratories, Inc. | RST | 100% (0-7 days), 100% (8-14 days), 100% (15+ days) | 95.9% | The ACON SARS-CoV-2 IgG/IgM Rapid Test is a lateral flow assay for the qualitative detection of IgM and IgG antibodies specific to SARS-CoV-2. The target antigen is not stated. The test results can be read by eye. Sensitivity was determined from 54 positive clinical samples, and specificity was determined from 73 negative clinical samples. Values presented here are for combined IgM/IgG. The test has also been independently validated by the NCI, with a combined sensitivity of 100%, and specificity of 96.2%. | Received EUA by the FDA
Innovita 2019-nCoV Ab Test (Colloidal Gold) | Innovita (Tangshan) Biological Technology Co., Ltd. | RST | (For combined IgG/IgM) 95.6% (8-14 days), 100% (15+ days) | 98.03% | The Innovita 2019-nCoV Ab Test (Colloidal Gold) is a lateral flow assay for the qualitative detection of IgM and IgG antibodies to SARS-CoV-2. The target antigens are recombinant S1 and NP proteins. The test may be read by eye. Sensitivity was determined through a multisite study, with 163 positive clinical samples and 305 negative clinical samples. Data presented here are from the inital multisite study, but a separate US study and longitudinal study were also performed (data available in IFU). Independent validation by the NCI showed a sensitivity of 100%, and a specificity of 97.5%. | Received EUA by the FDA
LYHER Novel Coronavirus (2019-nCoV) IgM/IgG Antibody Combo Test Kit (Colloidal Gold) | Hangzhou Laihe Biotech Co., Ltd | RST | 100% (IgM, 0-6 days), 0% (IgG, 0-6 days); 85.7% (IgM, 7-14 days), 76% (IgG, 7-14 days); 99.25% (IgM, 14+ days), 98.5% (IgG, 14+ days) | 99.43% | The LYHER Novel Coronavirus (2019-nCoV) IgM/IgG Antibody Combo Test Kit (Colloidal Gold) is a lateral flow assay for qualitative detection of antibodies specific to SARS-CoV-2. The target antigen is the S1 spike protein. Results are available in 10 minutes, but should not be read after 15 minutes. Sensitivity was determined from 178 positive clinical samples, and specificity from 349 negative clinical samples. The test was independently validated by the NCI, with a combined sensitivity of 100%, and specificity of 98.8%. | Received EUA by the FDA
(ARCHITECT and Alinity i systems) | Abbott Laboratories Inc. | Chemiluminescent microparticle immunoassay (CMIA) | 43% (3-7 days); 78.9% (8-14 days); 100% (15+ days) | 99.6% | The SARS-CoV-2 IgG assay detects IgG by measuring antibody-antigen binding in terms of relative light units (RLU). It is intended to be used with the ARCHITECT i1000SR and i2000SR systems, and the Alinity i systems. The December 2020 update includes the addition of this language, "“Due to the risk of false positive results, confirmation of positive results should be considered – using a second, different antibody assay that detects the same type of antibodies." Sensitivity was determined from 122 positive clinical samples, and specificity from 1070 clinical samples. | Received EUA by the FDA
Elecsys Anti-SARS-CoV-2 S (semi-quantitative) | Roche | EChLIA | 90.6% (0-7 days), 87% (8-14 days), 96.6% (15+days) | 99.98% | The Roche Elecsys Anti-SARS-CoV-2 S assay detects total antibodies (IgG, IgM, and IgA) against the SARS-CoV-2 spike protein RBD region. The test must be run on the Roche cobas e analyzer platform. The test takes approximately 18 minutes. They also tested the assay against 100 samples with antibodies to common human coronaviruses, with no cross-reactivity. Sensitivity was determined using 1485 positive clinical samples. Specificity was determined using 5991 negative samples. The IFU has detailed data on results from clinical performance testing. | Received EUA, CE mark
COVID-SeroKlir, Kantaro Semi-Quantitative SARS-CoV-2 IgG Antibody Kit | Kantaro Biosciences LLC | ELISA | 100% (0-7 days), 100% (7-14 days), 98.87% (15+ days) | 99.6% | The COVID-SeroKlir, Kantaro Semi-Quantitative SARS-CoV-2 IgG Antibody Kit is an assay to detect antibodies to SARS-COV-2 spike protein. The first ELISA detects antibodies against the spike RBD region, and the second step is a confirmatory second ELISA that confirms antibodies through reaction to full length spike protein. Samples should only be tested from individuals that are 15 days or more post symptom onset. Results can be qualitative or semi-quantitative. Sensitivity was determined from 251 positive clinical samples, of which 235 had days post symptom onset available. Specificity was determined from 284 negative samples. | Received EUA
Orawell IgM/IgG Rapid Test | Jiangsu Well Biotech Co., Ltd. | RST | 97.1% | 100% | The Orawell IgM/IgG Rapid Test is a lateral flow immunoassay intended for qualitative detection and differentiation of IgM and IgG antibodies to SARS-CoV-2 in human serum and acid citrate dextrose (ACD) plasma. The test was also validated by the NCI, with a sensitivity of 100%, and a specificity of 94.8%. The November updated adds an additional authorized distributor, Phase Scientific International Limited. | Received FDA EUA
BIOTIME SARS-CoV-2 IgG/IgM Rapid Qualitative Test | Xiamen Biotime Biotechnology Co., Ltd | RST | IgG: 46.38% (days 0-7), 67.65% (days 7-14), 100% (days 14+); IgM: 55% (days 0-7), 94% (days 7-14), 100% (days 14+) | 99.23%% | The November EUA updates adds an authorized distributor of the test, Telepoint Medical Services, LLC. This is a lateral flow assay that detects IgM and IgG against SARS-CoV-2 (specific antigen not stated). This has only been authorized for use with potassium EDTA plasma, serum or potassium EDTA venous whole blood. In testing for cross-reactivity with other samples from patients with non-SARS-CoV-2 infections, there was no cross-reactivity found. The EUA revision included changing the result reading time from 10 minutes to 20 minutes and replace the time after which results should not be read from 20 minutes to 30 minutes. They provide sensitivity and specificity stratified by time post symptom onset (presented here) and time post positive PCR. Sensitivity was determined from 120 positive clinical samples, and specificity from 315 clinical samples. The manufacturer also did a longitudinal study to determine time of seroconversion. The test was also validated by the NCI, with a combined sensitivity of 100% and specificity of 96.2%. The second EUA revision makes minor changes to the language surrounding the implications of a positive or negative result in the IFU. | Received EUA
Access SARS-CoV-2 IgM | Beckman Coulter | ChLIA | 91.7%(8-14 days); 98.3%(15+ days) | 99.9% | The Access SARS-CoV-2 IgM assay detects IgM to the RBD of the spike protein of SARS-CoV-2 . The test should only be run on samples 8 days to 30 days post symptom onset. Sensitivity was determined using 173 positive clinical samples, and specificity using 1400 negative samples. Sensitivity is presented here for samples 8 days or more post symptom onset. | Received FDA EUA, CE mark
cPass SARS-CoV-2 Neutralization Antibody Detection Kit | GenScript USA, Inc. | Modified ELISA | 100% | 100% | The cPass SARS-CoV-2 Neutralization Antibody Detection Kit is a modified ELISA that detects neutralizing antibodies that can block a recombinant RBD interaction with human ACE2. The kit was validated against a PRNT, using true SARS-CoV-2, for both PRNT50 and PRNT90. Sensitivity was using 26 positive samples (detected by PRNT) and 88 negative samples. There is cross reactivity with SARS-CoV-1 antibodies. | Received EUA
Dimension Vista SARS-CoV-2 Total antibody assay (COV2T) | Siemens Healthcare Diagnostics Inc. | ChLIA | 91.3% (7-13 days), 100% (15+ days) | 99.8% | The Dimension Vista COV2T assay is a chemiluminescent immunoassay used for the detection of total antibodies to SARS‑CoV‑2 in human serum and plasma from patients who may have been exposed to coronavirus disease (COVID‑19). The Dimension Vista COV2T assay is a homogeneous, sandwich chemiluminescent immunoassay based on LOCI technology. The first EUA revision makes several technical changes/additions to the IFU including clarifying reagent shelf stability, updating precision data, and allowing use of additional software/devices. The second EUA revision makes minor changes to the language surrounding the implications of a positive or negative result in the IFU. | Received EUA, available for purchase by healthcare professionals and researchers.
Dimension EXL SARS-CoV-2 Total antibody assay (CV2T) | Siemens Healthcare Diagnostics Inc. | ChLIA | 91.3% (7-13 days), 100% (15+ days) | 99.8% | The Dimension EXL CV2T assay is a chemiluminescent immunoassay used for the detection of total antibodies to SARS‑CoV‑2 in human serum and plasma from patients who may have been exposed to coronavirus disease (COVID‑19). The Dimension EXL CV2T assay is a homogeneous, sandwich chemiluminescent immunoassay based on LOCI technology. The LOCI reagents include two synthetic bead reagents and a biotinylated S1 RBD antigen. The first EUA revision makes several technical changes/additions to the IFU including clarifying reagent shelf stability, updating precision data, and allowing use of additional software/devices. The second EUA revision makes minor changes to the language surrounding the implications of a positive or negative result in the IFU. | Received EUA, available for purchase by healthcare professionals and researchers.
CareStart COVID-19 IgM/IgG | Access Bio, Inc. | RST | 96.7% (IgG); 89% (IgM); 98.44% (combined) | 99.45% (IgM); 99.45%(IgG); 98.9% (combined) | This is a lateral flow assay that detects IgM and IgG to recombinant nucelocapsid and spike RBD antigens of SARS-CoV-2. Sensitivity was determined from 64 positive clinical samples, and specificity from 182 negative samples. The company also stratifies by time post symptom onset, though most samples were from 15 days onward. This test has also been validated by the NCI, and the combined (IgG/IgM) sensitivity is 100%, and specificity is 97.5%. The EUA revision includes modifications to the product brand names which include "AQGene COVID-19 MDx RT-PCR" and "Rapid Response Liberty COVID-19 IgG/IgM." Other minor adjustments in the IFU, including the addtion of more external controls, are also detailed in the revision. As of April 2021, the test should not be used to measure vaccine responses, and that performance of the test has not been established in individuals that have received a COVID-19 vaccine. They also state that the serology test has not been validated for all variants circulating. | Received EUA
Platelia SARS-CoV-2 Total Ab assay | Bio-Rad | Modified ELISA | 98% | 99% | The October update includes language on false positive results: "a positive result may not indicate previous SARS-CoV-2 infection. Consider other information including clinical history and local disease prevalence, in assessing the need for a second but different serology test to confirm an immune response." The Platelia SARS-CoV-2 Total Ab assay measures IgM, IgG, and IgA antibodies to SARS-CoV-2. The target antigen is recombinant nucleocapsid protein, coupled with peroxidase that provides the read-out. Human serum or plasma samples are required.The sensitivity and specificity are not yet known, though it appears the readout of the assay may be quantiative or qualitative. | Received EUA
Q-Plex SARS-CoV-2 Human IgG (4 Plex) | Quansys Biosciences, Inc. | ELISA | 100% (0-7 days, n=4), 100% (8-14 days, n=13), 95.2% (15+days, n=22) | 99.7% | The Quansys Q-Plex SARS-CoV-2 Human IgG is a multiplexed assay to detect IgG in patient samples. The target antigens are the S1 and S2 subunits of SARS-CoV-2 spike protein. The test should be run on the Quansys Q-View Imager LS or Quansys Q-View Imager Pro and Q-view Software platforms. Sensitivity was determined using 39 positive clinical samples (stratified by time post symptom onset), and specificity using 585 negative samples. | Received EUA
VITROS Immunodiagnostic Products Anti-SARS-CoV-2 IgG Reagent Pack | Ortho-Clinical Diagnostics, Inc. | ChLIA | 83% (12-15 days), 100% (15 days+) |  | The October 2020 update included language to account for false positives, in that a positive test does not necessarily indicate previous infection (especially in lower prevalence areas). In addition, the updated includes "“Due to the risk of false positive results, confirmation of positive results should be considered – using a second, different antibody assay that detects the same type of antibodies." The VITROS Immunodiagnostic Products Anti-SARS-CoV-2 IgG Reagent Pack detects antibodies specific to SARS-CoV-2 and is designed to be used with VITROS Immunodiagnostic Products and the VITROS ECi/ECiQ/3600 Immunodiagnostic System and the VITROS 5600/XT 7600 Integrated Systems. Sensitivity was approximately 87.5% and specificity is approximately 100%. | Received EUA by the FDA
Elecsys Anti-SARS-CoV-2 | Roche | EChLIA | From 0-6 days, 65.5%; from 7-13 days, 88.1%; from 14 days onward, 100% | 99.81% | The October update includes language on false positive results: "“A positive result may not indicate previous SARS-CoV-2 infection. Consider other information including clinical history and local disease prevalence, in assessing the need for a second but different serology test to confirm an immune response." The test couples chemiluminescence immunoassay technology with an electric pulse that allows for rapid quantifiation of captured antigen-antibody complexes from patient serum samples. The target antigen is the viral N protein.The test assays for IgM and IgG, and takes about 18 minutes. The test should be run on the Elecsys Anti-SARS-CoV-2 on the cobas e 411, cobas e 601, cobas e 602, or cobas e 801 analyzers. Sensitivity was determined from 204 samples, and specificity from 5,272 negative samples. | Received EUA, available for purchase by healthcare professionals and researchers. | https://www.fda.gov/media/137602/download
LIAISON SARS-CoV-2 S1/S2 IgG | DiaSorin Inc. | ChLIA | 90% (8-14 days), 97% (15+ days) | 98% | The October 2020 update included language to account for false positives, in that a positive test does not necessarily indicate previous infection. In addition, the updated includes "“Due to the risk of false positive results, confirmation of positive results should be considered – using a second, different antibody assay that detects the same type of antibodies." This test is designed to run on the LIAISON XL analyzer and detects IgG specific to recombinant spike protein subunits 1 and 2 (S1 and S2)." Specificity was approximately 98%. Sensitivity was approximately 90% on samples obtained 5-15 days post-symptom onset. For samples taken greater than 15 days post-symptom onset, the sensitivity was approximately 97%. This test is designed to run on the LIAISON XL analyzer and detects IgG specific to recombinant spike protein subunits 1 and 2 (S1 and S2). Specificity was approximately 98%. Sensitivity was approximately 90% on samples obtained 5-15 days post-symptom onset. For samples taken greater than 15 days post-symptom onset, the sensitivity was approximately 97%. | Received EUA by the FDA
Biohit SARS-CoV-2 IgM/IgG Antibody Test Kit | Biohit Healthcare (Heifei) Co. Ltd. | RST | 33% (IgM, days 1-7), 56.6% (IgG days 8-14), 83.% (IgM days 8-14), 96.2% (IgG days 15+), 97.7% (IgM days 15+) | 99.46% (IgM), 100% (IgG) | The October 2020 update includes updating the name of the kit to the ThermoGenesis SARS-CoV-2 IgM/IgG Antibody Test Kit. In addition, the expiration date was extended to 6 months (5 months for the controls). This test detects IgM and IgG through immunchromatography. It is not intended for use early in infection, with optimal performance after 15 days post symptom onset. The target antigen is recombinant N-protein. Specificity was determined from 186 negative samples (including 78 samples positive for other respiratory diseases), and sensitivity was determined from 197 serum samples. | Received EUA by the FDA
MosaiQ COVID-19 Antibody Magazine | Quotient Suisse SA | solid-phase photometric immunoassay | 100% | 99.8% | The MosaiQ COVID-19 Antibody Magazine detects IgM and IgG in patient serum samples, and operates using a microarray to detect antibodies. The test must be run on the MosaiQ 125 instrument. The target antigen is the Spike protein. The test may detect antibodies as soon as 7 days post-PCR positivity. Sensitivity was determined using 30 positive clinical samples (8-14 days post PCR positivity (n=15), and ≥ 15 days (n=15)), and specificity using 400 negative samples. An independent validation study did include samples from 0-7 days post PCR positivity (n=5), 8-14 days (n=9), and 15+days (n=86). Sensitivity was 100%, except after 15+ days, where it was 93%. | Received EUA, and CE mark
AdviseDx SARS-CoV-2 IgM (ARCHITECT and Alinity i systems) | Abbott Laboratories Inc. | CMIA | 95% | 99.56% | The AdviseDx SARS-CoV-2 IgM test can be run on the Alinity i or Architect platforms, with separate instructions for use for each. This test should only be run with samples 15 days to 30 days post symptom onset. Sensitivity presented here is only for 15 days or more post-symptom onset. Sensitivity was determined using 40 positive clinical samples (for each platform), and specificity was determined using 2965 (Architect) and 2985 (Alinity) negative samples. | Received FDA EUA
Maverick SARS-CoV-2 Multi-Antigen Serology Panel v2 | Genalyte, Inc | photonic ring immunoassay (PRI) | 66.7% (0-7 days); 90.91% (8-14 days); 96.1% (15+ days) | 97.7%% | The Maverick SARS-CoV-2 Multi-Antigen Serology Panel v2 is a qualitative assay to detect IgG and IgM antibodies specific to SARS-CoV-2. The photonic ring immunoassay uses silicon photonics to multiplex several antigens on a small chip. This allows for simultaneous detection of a variety of antibodies to different viruses. This assay can detect antibodies to SARS-CoV-2 nucleocapsid, full length spike, S1, S2, and RBD subunits. It also detects antibodies to the spike protein of the common coronaviruses, 229E, OC43, and HKU1, and to the nucleoprotein of NL63. Further, it can detect antibodies to influenza A H1 and H3 proteins, the S1 subunit of MERS, and the nucleoprotein of SARS-CoV-1. However, the test only reports positivity/negativity for SARS-CoV-2. The test must be run on the Maverick platform, and a proprietary algorithm analyzes the signal to determine test results. This table depicts the sensitivity and specificity measures for SARS-CoV-2 only. Sensitivity was determined using 338 positive clinical samples, stratified by time post sympto onset. Specificity was determined from 862 negative samples. | Received FDA EUA
ZEUS ELISA SARS-CoV-2 IgG Test System | ZEUS Scientific, Inc. | ELISA | 100% | 99.1% | The ZEUS ELISA SARS-CoV-2 IgG Test System is an indirect, antibody capture enzyme-linked immunosorbent assay designed to detect IgG class antibodies to SARSCoV-2 (novel 2019 Coronavirus) in human serum or plasma collected in CLIA certified laboratories. The wells of the plastic microwell strips are coated with a mixture of recombinant S1 receptor binding domain (RBD) viral protein and recombinant nucleoprotein as the antibody capture antigens. Sensitivity and specificity were determined from 28 positive and 214 negative samples. | Received FDA EUA
OmniPATH COVID-19 Total Antibody ELISA Test | Thermo Fisher Scientific | ELISA | 19% (0-7 days), 76.7% (8-14 days), 100% (15+days) | 100% | The OmniPATH COVID-19 Total Antibody ELISA Test detects IgG, IgM, and IgA antibodies. The ELISA can be run manually, or using the Dynex AGILITY automated ELISA workstation.The target antigen is recominant spike protein. Sensitivity was measured using 91 positive clinical samples, stratified by time post symptom onset. Specificity was measured using 162 negative samples. The test was also validated by the NCI, finding a sensitivity of 96.7% and specificity of 97.5%. | Received FDA EUA and CE mark
Nirmidas Biotech, Inc. | Nirmidas Biotech, Inc. | RST | IgM: 97.1% IgG: 100% | 94.8% | This test uses the principle of colloidal gold immunochromatography, and capture method to detect COVID-19 (SARS-CoV-2) IgM and IgG antibodies in human serum and plasma. When the sample contains COVID-19 (SARS-CoV-2) IgM / IgG antibody and the concentration is greater than or equal to the minimum detection limit, the antibody binds to the S1 and RBD subunits of spike protein immobilized in test region M / test region G and captured by the secondary antibody to produce a red reaction line. | Received FDA EUA
FREND COVID-19 total Ab | NanoEntek America, Inc. | immunofluorescence assay | 100% | 100% | The FREND™ cartridge utilizes micro-fluidics lateral flow technology where the analyte of interest in the sample forms immune complexes while moving through the fluidics pathway in the cartridge. A specimen is added to sample dilution tube and mixed. A well-mixed sample of 35 μL is transferred to the sample inlet of a single use FREND™ COVID-19 total Ab cartridge. The cartridge is then placed into the FREND™ System, which is programmed to begin analysis once the sample has reacted with the reagents. The FREND™ System is a bench-top fluorescence reader containing a touchscreen user interface. Sensitivity and specificity was determined with 33 positive and 12 negative samples. The test was also independently validated by the NCI with a sensitivity of 96.7% and specificity of 98.8%. | Received FDA EUA and CE mark
DiaSorin LIAISON SARS-CoV-2 IgM Assay | DiaSorin, Inc. | ChLIA | 92.6% | 99.3% | The method for qualitative determination of specific IgM to SARS-CoV-2 is an indirect chemiluminescence immunoassay (CLIA). SARS-CoV-2 spike receptor binding domain (RBD) antigen is used for coating magnetic particles (solid phase). The light signal, and hence the amount of isoluminol-antibody onjugate, is measured by a photomultiplier as relative light units (RLU) and indicates the presence or absence of antibodies to SARS-CoV2 present in calibrators, samples or controls. Sensitivity and specificity was determined with 81 positive and 2473 negative samples. | Received FDA EUA
Assure COVID-19 IgG/IgM Rapid Test Device | Assure Biotech (Hangzhou Co, Ltd.) | RST | 95.1% | 100% | This test's EUA was updated on September 23,2020. The updates include that the test can now be used with fingerstick samples, and in point-of-care settings.This lateral flow assay tests for both IgM and IgG in patient serum. The target antigens are recombinant Nucleocapsid protein and Spike protein (S1 region). Sensitivity was determined from 61 positive clinical samples, and specificity from 105 negative samples. The company also provides sensitivity stratified by time post symptom onset, though the overall sensitivity is presented here. Overall, the kit had highest sensitivity with venous whole blood. The kit has also been independently validated by the NCI, as presented by the manufacturer. Combined IgG/IgM sensitivity was 100%, and specificity was 98.8%. | Received EUA
SARS-CoV-2 IgG (ARCHITECT and Alinity i) | Abbott Laboratories Inc. | Chemiluminescent microparticle immunoassay (CMIA) | 0% (0-3 days), 25% (3-7 days), 86.4% (8-13 days), and 100% (14+ days) | 99.63% | This is an update to the EUA received 4/26/2020, for the ARCHITECT i SARS-CoV-2 IgG and Alinity i SARS-CoV-2 IgG assays. This test detects IgG by measuring antibody-antigen binding in terms of relative light units (RLU). It is intended to be used with the ARCHITECT i1000SR and i2000SR systems, as well as the Abbott Alinity i system. Sensitivity and specificity were determined on the Architect system, with 122 positive clinical samples and 1070 negative samples (including non-COVID respiratory illnesses). Sensitivity was stratified by time post symptom onset. | Received EUA by the FDA (most recently revised on 9/18/2020)
MAGLUMI 2019-nCoV IgM/IgG | Shenzhen New Industries Biomedical Engineering Co., Ltd. (SNIBE) | ChLIA | 43.8% (0-7 days), 93.4% (8-14 days), 100% (15+ days) | 98.67% | The MAGLUMI 2019-nCoV IgM/IgG assay uses a capture chemiluminescent immunoassay to detect IgM and IgG antibodies to SARS-CoV-2 recombinant spike and nucleocapsid proteins. This test should be run on the MAGLUMI 2000 platform. IgM and IgG are measured in two separate cassettes, though the results are displayed in a single report. Sensitivity was determined using 264 positive clinical samples, stratified by time. The manufacturer provides separate IgG and IgM sensitivity, though the combined sensitivity is presented here. Specificity was determined from 226 negative samples. | Received EUA, and CE mark
BioCheck SARS-CoV-2 IgM Antibody Test Kit | BioCheck, Inc. | ChLIA | 93.7%(8-14 days), 88.9% (15+ days) | 97.2% | The BioCheck SARS-CoV-2 IgM Antibody Test Kit uses an indirect method to determine the presence of SARS-CoV-2 IgM antibodies in a human serum. In the cassette, the SARS-CoV-2 IgM antibodies in a patient serum sample bind to biotinylated SARS-CoV-2 S1 antigens. Excess streptavidin coated magnetic beads bind with high affinity to biotinylated SARS-CoV-2 S1 antigen to form complexes. A luminescent substrate is added, and the enzymes in the complexes catalyze the luminescent substrate to form unstable excited state intermediates. When the excited intermediates return to the ground state, photons are emitted. The number of photons produced is positively correlated with the level of SARS-CoV-2 IgM antibodies in the sample. | Received FDA EUA and CE mark
BioCheck SARS-CoV-2 IgG Antibody Test Kit | BioCheck, Inc. | ChLIA | 100% | 100% | The BioCheck SARS-CoV-2 IgG Antibody Test Kit uses an indirect method to determine the presence of SARS-CoV-2 IgG antibodies in a human serum. In the cassette, the SARS-CoV-2 IgG antibodies in a patient serum sample bind to biotinylated SARS-CoV-2 S1 antigens. Excess streptavidin coated magnetic beads bind with high affinity to biotinylated SARS-CoV-2 S1 antigen to form complexes. A luminescent substrate is added, and the enzymes in the complexes catalyze the luminescent substrate to form unstable excited state intermediates. When the excited intermediates return to the ground state, photons are emitted. The number of photons produced is positively correlated with the level of SARS-CoV-2 IgG antibodies in the sample. | Received FDA EUA and CE mark
COVID-19 ELISA pan-Ig Antibody Test | University of Arizona Genetics Core for Clinical Services | ELISA | 97.5% | 99.1% | The COVID-19 ELISA pan-Ig Antibody Test is an Enzyme-Linked Immunosorbent Assay (ELISA) intended for the qualitative detection of total antibodies (including IgA, IgG, and IgM) to SARS-CoV-2 in human serum. Sensitivity and specificity was determined from 40 positive samples and 320 negative samples. | Received FDA EUA
TBG SARS-CoV-2 IgG / IgM Rapid Test Kit | TBG Biotechnology Corp. | RST | IgG: 96.4% IgM: 87.5% | 99.8% | The TBG SARS-CoV-2 IgG / IgM Rapid Test Kit is a lateral flow immunoassay intended for qualitative detection and differentiation of IgG and IgM antibodies to SARS-CoV-2 in human serum and plasma. Manufacterer reported sensitivity and specificity was determined from 56 positive samples and 482 negative samples. The test was also validated by the NCI, finding a sensitivity of IgG 93.3%/ IgM 93.3% and specificity of IgG 96.2%/ IgM 95.0%. | Received FDA EUA and CE mark, NCI validation studies complete
Tell Me Fast Novel Coronavirus (COVID-19) IgG/IgM Antibody Test | Biocan Diagnostics Inc. | RST | IgG: 93.3% IgM: 90% | IgG: 96.2% IgM: 98.7% | Biocan Tell Me Fast Novel Coronavirus (COVID-19) IgG/IgM Antibody Rapid Test is a lateral flow chromatographic immunoassay. The T2 line is pre-coated with monoclonal anti-human IgG for the detection of IgG anti-SARS-CoV-2, the T1 line is pre-coated with reagents for the detection of IgM anti-SARS-CoV-2. The sensitivity was determined from 30 positive clinical samples and specificity from 79 negative samples. | Received FDA EUA and CE mark
BioCheck SARS-CoV-2 IgG and IgM Combo Test | BioCheck, Inc. | ChLIA | 100% (0-7 days, IgM); 93.8% (7-14 days, IgM); 88.9% (15+days, IgM); 100% (0-7 days, IgG); 100%(7-14 days, IgG); 100% (15+days, IgG) | 97.2% (IgM), 100% (IgG) | The BioCheck SARS-CoV-2 IgG and IgM Combo Test is a CLIA based assay that detects antibodies specific to the S1 region of spike protein. The assay, which uses special casettes, should be run on the MS-Fast Automated Chemiluminescent Immunoassay Analyzing System. Sensitivity was determined from 110 positive clinical samples, and specificity from 143 negative samples. Sensitivity was further stratified by time post symptom onset. | Received EUA and CE mark
Diazyme DZ-Lite SARS-CoV-2 IgM CLIA Kit | Diazyme Laboratories, Inc. | ChLIA | 26.1% (0-7 days); 83.3% (7-14 days), 94.4% (15+days) | 98.3% | The Diazyme DZ-Lite SARS-CoV-2 IgM CLIA Kit detects IgM specific to SARS-CoV-2. The test uses a recombinant virus antigen, though it is not stated which protein is used. The manufacturer recommends that the kit be used at least 15 days post symptom onset. Sensitivity was determined from 155 positive clinical samples, and specificity from 302 negative samples. Sensitivity was stratified by time post symptom onset (108 samples were from after 15 days). | Received EUA
VIDAS SARS-CoV-2 IgM | bioMérieux SA | Enzyme linked fluorescent assay (ELFA) | 53.8% (0-7 days), 100% (7-14 days), 100%(15+) | 99.4%% | The VIDAS SARS-CoV2 IgM measures IgM antibody to the S1 unit of the spike protein. The system is completely automated, and provides a fluorescent readout that correlates to titers of antibody present in the patient sample. Sensitivity was determined from 111 positive clinical samples, stratified by time, and 308 negative clinical samples. Only 20 samples were from after 7 days post symptom onset. 91 were from 0-7 days. | Received EUA, has CE mark.
VIDAS SARS-CoV-2 IgG | bioMérieux SA | Enzyme linked fluorescent assay (ELFA) | 49% (0-7 days), 100% (7-14 days), 100% (15+days) | 99.9%% | The VIDAS SARS-CoV-2 IgG measures IgG antibody to the S1 unit of the spike protein. This system is completely automated, and provides a fluorescent readout that correlates to titers of antibody present in the patient sample. Sensitivity was determined from 120 positive clinical samples, and specificity from 989 negative samples. The positive samples were stratified by time since symptom onset, and only 17 were from 15 days onward. | Received EUA, has CE mark.
WANTAI SARS-CoV-2 Ab ELISA | Beijing Wantai Biological Pharmacy Enterprise Co., Ltd. | ELISA | 55.4% (0-7 days), 84.8% (7-14 days), 98.7% (15+ days) | 98.6% | The WANTAI SARS-CoV-2 Ab ELISA test is an ELISA based assay, detecting IgM and IgG reactive to the spike protein receptor binding domain. Sensitivity was determined from 235 positive samples, and specificity from 501 negative samples stratified by time since symptom onset. The test was also validated by the NCI, finding a sensitivity of 96.7% and specificity of 97.5%. | Received EUA
Atellica IM SARS-CoV-2 IgG (COV2G) | Siemens Healthcare Diagnostics Inc. | ChLIA | 56% (0-6 days), 92%(7-14 days), 100% (14+ days) | 99.9% | This is a chemiluminescent immunoassay to detect IgG in patient samples, using the Atellica platform. The target antigen is the RBD of the S1 domain of the spike protein. This assay provides a semi-quantitative (meaning an estimate/relative amount of antibodies) readout within about 25 minutes. Sensitivity and specificity were determined using 197 positive clinical samples and 1841 negative samples. Both measures were stratified by time post symptom onset. | Received EUA
ADVIA Centaur SARS-CoV-2 IgG (COV2G) | Siemens Healthcare Diagnostics Inc. | ChLIA | 53.5% (0-6 days), 93.4% (7-14 days), 100% (14+ days) | 99.9% | This is a chemiluminescent immunoassay to detect IgG in patient samples, using the ADVIA Centaur XP or XPT platforms. The target antigen is the RBD of the S1 domain of the spike protein. This assay provides a semi-quantitative (meaning an estimate/relative amount of antibodies) readout within about 25 minutes. Sensitivity and specificity were determined using 189 positive clinical samples and 1831 negative samples. Both measures were stratified by time post symptom onset. | Received EUA
xMAP SARS-CoV-2 Multi-Antigen IgG Assay | Luminex | Fluorescent microbead-based immunoassay (FMIA) | Serum: 71% (days 0-7), 71.4% (8-14 days), 96.2% (14+ days); Plasma: 100% (0-7 days), 82% (8-14 days), 96.6% (14+ days) | Serum: 100%, Plasma: 99.2% | This is a multiplexed, bead based assay that tests for antibodies against the nucleocapsid protein, the receptor binding domain of spike, and the S1 subunit of spike protein. The test should be run using the Luminex® xMAP technology and works with Luminex 200, MAGPIX, and FLEXMAP 3D System. This test is specific for IgG. Sensitivity was determined from 112 serum samples, stratified by time since symptom onset. Specificity was from 309 negative serum samples. They also tested plasma sensitivity, stratified by time (n=46). They validated sensitivity and specificity on all platforms suitable for the test. | Received FDA EUA
Sienna-Clarity COVIBLOCK COVID-19 IgG/IgM Rapid Test Cassette | Salofa Oy | RST | 100% for IgM and IgG | 100% for IgM and IgG | The January 2021 update includes the addition of fingerstick whole blood sample use. Fingerstick whole blood can be used in CLIA certified labs, or in POC settings operating under a CLIA waiver, Certificate of Compliance, or Certificate of Accreditation. The SiennaTM-Clarity COVIBLOCKTM COVID-19 IgG/IgM Rapid Test Cassette is a membrane-based lateral flow immunoassay intended for the qualitative detection and differentiation of IgG and IgM antibodies to SARS-CoV-2 in human serum, plasma (dipotassium EDTA, sodium citrate, sodium heparin) and venous whole blood (dipotassium EDTA, sodium citrate, sodium heparin). NCI cross-validation studies reported sensitivity of 90% for IgM and 93.3% for IgG and specificity of 100% for IgM and 98.8% for IgG. | Received FDA EUA
WANTAI SARS-CoV-2 Ab Rapid Test | Beijing Wantai Biological Pharmacy Enterprise Co., Ltd. | RST | 94.7% | 98.9% | This is a lateral flow assay that detects total antibody to the RBD region of the spike protein of SARS-CoV-2. Sensitivity was determined from 132 positive clinical cases, and specificity from 271 negative cases. They also post sensitivity and specificity stratified by time since symptom onset.The test has been independently validated by several institutions, including the NCI which found a sensitivity of 100% and specificity of 98.8%. | Received EUA and CE mark.
Diazyme DZ-Lite SARS-CoV-2 IgG CLIA Kit | Diazyme Laboratories, Inc. | ChLIA | 91.7% | 97.4% | The Diazyme DZ-Lite SARS-CoV-2 IgG CLIA Kit is an indirect chemiluminescence immunoassay. The light signal is measured by the DZ-lite 3000 Plus Chemiluminescence Analyzer as relative light units (RLUs), which is evaluated against the established cutoff to determine the final result. The product contains a Low and a High Calibrator to execute calibration operation. | Received EUA
Access SARS-CoV-2 IgG | Beckman Coulter, Inc. | ChLIA | 75% (0-7 days), 95.3% (8-14 days), 96.8% (15+ days) | 99.6% | This is a two step CLIA that uses the target antigen of the receptor binding domain of S1 of the spike protein. It detects IgG in patient samples. Sensitivity was determined using 192 positive clinical samples, stratified by time since symptom onset. Specificity was determined from 1400 samples collected before December 2019. | Received EUA by the FDA
Babson Diagnostics aC19G1 | Babson Diagnostics | ChLIA/indirect sandwich immunoassay | 66.7% (8-14 days), 100% (15+ days), 100% (asymptomatic) | 100% | This test is a CLIA based assay that detects IgG antibodies. The target antigen is the S1 region of the spike protein. The test should be run on the Siemens Atellica IM platform. Sensitivity was determined from 34 positive samples, stratified by time post symptom onset and 2 of which were asymptomatic individuals. Specificity was determined from 100 individuals. | Received EUA by the FDA
SARS-CoV-2 RBD IgG test | Emory University | ELISA | 73% (days 0-7), 100% (days 8-14), 100% (days 14+) | 97.7% (prior to pandemic), 94.4% (PCR negative HCW) | This test detects IgG antibodies specific to a recombinant receptor binding domain (RBD) of the spike protein. It should be run on the BioTek 800 TS Absorbance Reader. Sensitivity was determined from 231 positive samples, and specificity from 388 negative samples--including samples taken prior to the pandemic and rRT-PCR negative health care workers during the pandemic. These results were stratified based on days post-positive rRT-PCR. | Received EUA by the FDA
SCoV-2 Detect IgG ELISA | InBios International, Inc. | ELISA | 97.8% | 98.9% | The SCoV-2 Detect™ IgG ELISA is a qualitative immunoassay for the detection of IgG antibodies targeting SCoV-2 related antigens. Diluted serum specimens are added to antigencoated wells and incubated. After incubation and washing, human antibodies targeting SARS-CoV-2 antigens remain bound to the plate surface. Secondary antibody conjugated to horseradish peroxidase (HRP) targeting human IgG is then added to each well. After incubation, the ELISA wells are washed once again before a tetramethylbenzidine (TMB) substrate is added. An acidic stopping solution is finally used to stop the reaction and the degree of enzymatic turnover of the substrate is determined by absorbance measurement at 450 nanometers. | Received EUA, available for purchase by healthcare professionals and researchers.
RightSign COVID-19 IgG/IgM Rapid Test Cassette | Hangzhou Biotest Biotech Co., Ltd | RST | 96.3% (8-14 days), 94.9% (15+ days) (IgM); 91.6% (8-14 days), 96.2% (15+ days) (IgG) | 98.1% (IgM), 99.52% (IgG) | This RST tests for IgM and IgG in patient whole blood, serum, or plasma. The target antigen is recombinant spike protein receptor binding domain. Results should only be read after 10 minutes of incubation, but before 20 minutes. Clinical samples used for internal validation were from 3 different sites in China. Sensitivity and specificity presented here are for overall measures, though the company provides information for days 1-7 post symptoms, 8-14, and days 15+. Sensitivity and specificity were determined from 320 positive and 210 negative samples. Sensitivity increases after day 8 post symptom onset. | Received EUA, available for purchase by healthcare professionals and researchers.
Vibrant COVID-19 Ab Assay | Vibrant America Clinical Labs | ChLIA | 98.1% | 98.6% | The Vibrant COVID-19 Ab assay is a chemiluminescence immunoassay (CLIA) intended for the qualitative detection and differentiation of IgM and IgG antibodies to SARS-CoV-2 in human serum or Dry Blood Spot (DBS) using fingerstick blood specimen collected by their health care provider. | Received EUA, available for purchase by healthcare professionals and researchers.
COVID-19 IgG/IgM Rapid Test Cassette (Whole Blood/Serum/Plasma) | Healgen Scientific LLC | RST | 96.7% (IgG),86.7 % (IgM), 96.7% combined | 98% (IgG), 99% (IgM), 97% combined | This is a lateral flow assay based test, that uses human venous whole blood, plasma from anticoagulated blood, or serum to detect IgG and IgM antibodies. The test takes approximately 10 minutes, and is a qualitative test for presence of antibodies. The antigen target is not stated. Sensitivity was determined from 90 positive clinical samples, and specificity from 101 negative clinical samples from two testing sites in China. The test was also independently validated, with sensitivity of 100% and specificity of 97.5%. | Received EUA and CE/IVD, available for purchase by healthcare professionals and researchers.
Atellica IM SARS-CoV-2 Total (COV2T) | Siemens Healthcare Diagnostics Inc. | ChLIA | 60.7% (0-6 days post PCR positive),97.5% (7-13 days post PCR positive), 100% (14+ days post PCR positive). | 99.82%% | This CLIA based assay uses the Atellica platform to detect total antibody to the S1 receptor binding domain. The test is fully automated, all-in-one step sandwich chemiluminescent immunoassay. It uses streptavidin-biotin based capture. Sensitivity was from 250 clinical positive samples, with times since PCR positivity listed. Specificity was determined from 1091 samples, including 98 pregnant women. | Received EUA, available for purchase by healthcare professionals and researchers.
ADVIA Centaur SARS-CoV-2 Total (COV2T) | Siemens Healthcare Diagnostics Inc. | ChLIA | 65.1% (0-6 days post PCR positive),97.5% (7-13 days post PCR positive), 100% (14+ days post PCR positive). | 99.81%% | This CLIA based assay uses the ADVIA Centaur platform to detect total antibody to the S1 receptor binding domain. The test is fully automated, all-in-one step sandwich chemiluminescent immunoassay. It uses streptavidin-biotin based capture. Sensitivity was determined using 262 clinical samples, with times listed post PCR postivity. Specificity was from 1589 people, including 100 pregnant women. | Received EUA, available for purchase by healthcare professionals and researchers.
Anti-SARS-CoV-2 ELISA (IgG) | Euroimmun AG | ELISA | From 0-10 days, 13.9%; from 11-20 days, 61.1%; from 21 days onward, 100%. 90% by NCI validation | 100% by NCI validation | The Euroimmun AG IgG ELISA assay uses the viral S1 region of the spike protein to quantify IgG antibodies present in serum samples. Sensitivity and specificity were determined compared to onset of symptoms in PCR positive and negative samples. Sample sizes varied over time. | Received EUA, available for purchase by healthcare professionals and researchers.
New York SARS-CoV Microsphere Immunoassay for Antibody Detection | Wadsworth Center, New York State Department of Health | Microsphere immunoassay | Not stated | 93-100% | This test assays for total antibodies (IgG, IgM, and IgA) in human serum samples. The test uses full-length recombinant nucleocapsid protein from SARS-CoV-1 (it is stated that the sequence is similar enough to use the other virus' N protein-90% homologous). The test should be run with a Luminex FlexMap dual laser cytometer. | Received EUA, available for purchase by healthcare professionals and researchers.
COVID-19 ELISA IgG Antibody Test | Mount Sinai Laboratory | ELISA |  |  | This test detects, qualitatively, IgG present in the serum of patients. The ELISA based method uses a 1:50 dilution of human serum that is flowed over a plate coated with the spike protein receptor binding domain (RBD). Sensitivity and specificity are not yet available. | Received EUA by the FDA
qSARS-CoV-2 IgG/IgM Rapid Test | Cellex Inc. | RST | 93.8% | 95.6% | This is a rapid serology test, lateral flow assay, which detects IgM and IgG to the nucelocapside protein of SARS-CoV-2. The sensitivity is 93.8% and specificity is 95.6%, when tested at 2 Chinese hospitals in a total of 128 COVID19 positive patients, and 250 COVID19 negative patients (as detected by RT-qPCR). | Approved by FDA for EUA on diagnostics, has CE approval